## Slide 1: US-guided visual servoing system for non-invasive therapeutic HIFU
Joonho Seo
Mitsuishi Sugita Laboratory
Dept. of Mechanical Engineering ,The University of Tokyo
GMSI 2010 Workshop Tour

## Slide 2: Therapeutic HIFU
High Intensity Focused Ultrasound
Non-invasive treatment by,
Measured 3D acoustic pressure distributions
Thermal ablation by HIFU
Pulverization by cavitations
JC HIFU System (in China)
Sonablate® 500 system
Commercial Therapeutic HIFU systems
MRgHIFU, Philips

## Slide 3: US-guided HIFU treatment
Target recognitions.
Monitoring.
Tracking respiratory movement.
And so on…
- Realize anesthetic treatment.
Reduce treatment time.
Safe treatment
Still has many problems…
Construction of US-guided HIFU therapy system for target movement by respiration
Motion compensation method
for point and volumetric targets
Goal
Solution

## Slide 4: Prototype system
HIFU Transducer and bi-plane US probes
600
1350
730
US raw signal data
Image Processing
Controller
Pulse
generator
Image conversion
Control signal (1kHz)
Slave Position
Amplifier
Pulse
Generator
Syncronization
XYZ
Stage
Transducer
Ultrasound
Probes
US-guided tracking system oveview
Real-time image processing
Tracking 3D position of stone

## Slide 5
Tracking model stone in extracted swine kidney
Image processing and calculation
Hardware
Software
Point/blob recognition and tracking
Experiment for “Tracking kidney stone”

## Slide 6: HIFU therapy for Kidney tumors, scenario
3D voxel image by MD/spiral CT
Renal cell carcinoma
Preoperative
Surface model with marked tumors and irradiation plans
Intra-operative
Target recognition by REGISTRATION of organ (US) and its 3D model.
HIFU transducer
?
Preoperative mesh model for organ

## Slide 7: Registration
Real-time intra-operative US image
Pre-operative 3D organ model with target
?

## Slide 8: Phantom experiment
B-Mode construction
Build spatial US image
Speckle noise reduction
Create seed contour
Align mesh to boundary pts.
Iteration
Pose estimation process
Registration
Preprocess
Boundary points search
Imaging by US machine
Linear interpolated raw RF data
3D Transform
Section normal
Feature points
Sections
US plane
Feature points
3D Mesh model
Registered
From phantom kidney
In US imaging system
Filtering outliers

## Slide 9: “Pose estimation by static biplane US and 3D preoperative model”
(1)
(2)
(3)
(4)
Chart: 
 | 3.58 | 3.53 | 1.18 | 1.17 | 1.41 | 1.35 | 1.15 | 0.95 | 0.8 | 0.71 | 0.66 | 0.68 | 0.83 | 0.78 | 0.55 | 0.53 | 0.76 | 0.49 | 0.49
Num. of iteration
Error(mm)
(1)
(2)
(3)
(4)

## Slide 10: Real-time calculation test
* Real time pose- estimation for a moving model along X axis
* Real time pose- estimation for a moving model along Y axis
Chart: 
| Grab | BConversionL | BConversionR | BoundaryExtL | BoundaryExtR | Registration
 | 15.4813 | 25.4786 | 25.4803 | 13.6471 | 13.2758 | 1.769
Total processing time = 56.17 ms, 17.8 fps Imaging = 40.96 ms
Boundary + registration = 15.21 ms
Time for a single calculation, in average…

## Slide 11: And then we will do…
Improve algorithm robustness for in-vivo experiment.
Accelerate calculations by GPGPU.
Real-time tracking by motion estimation methods.
HIFU irradiation plan for volumetric targets.
Consider organ deformations.
And so on…

## Slide 12: Thank you!
Questions or Suggestions ?